## Inputs
Task 1.2 uncertainty budget: scenario inputs
Input | Value | Unit | Source / assumption
Starting OD, C_0 | 1.5 | OD | Scenario choice: a typical well in the observed 1.0-1.8 range (user's plates). Must be >= C_target.
Target OD, C_target | 1 | OD | Process setpoint (assignment section 1.1).
Nominal transfer volume, V~_transfer | 100 | uL | Half of the 200 uL source well (assignment section 1.1).
Tip size (largest dispense) | 20 | uL | 20 uL tip on the Flex 1-50 uL pipette (assignment section 1.1). Use 20 or 50. For the 50 uL tip, set the CV (B8) to 0.07 (50 uL tip, 1 uL row); the bias bound stays 0.08.
Random error, CV | 0.035 | fraction | Opentrons spec, 20 uL tip, 1 uL row: Random Error 3.50 % CV. A CV is one standard deviation, so u = CV (GUM 4.3.3). Applied to every dispense (20 uL tip). For the 50 uL tip use 0.07.
Bias bound, a (half-width) | 0.08 | fraction | Opentrons spec, 20 uL tip, 1 uL row: Systematic Error +/-8.00 %. Rectangular distribution, u = a/sqrt(3) (GUM 4.3.7).
Correlation of transfer and dilution biases, r | 1 | - | Assumption 3 (project document): the bias is a property of the pipette, shared by every dispense of a run, so r = +1 (GUM 5.2.2 eq. 14, F.1.2.3). Lower r to test sensitivity.
Coverage factor for expanded uncertainty | 2 | - | U = k u_c (GUM 6.2.1).
Tolerance half-width | 0.02 | OD | Normalization range 1.0 +/- 0.02 OD (assignment section 1.1).

## Budget
Task 1.2 uncertainty budget (GUM law of propagation, eq. 13)
Derived quantities
Nominal media volume, V~_d = V~_transfer (C_0/C_target - 1) | 50 | uL | Controller's dose from the nominal mass balance (general-form note, layer 3).
Media fraction, 1 - C_target/C_0 | 0.33333 | - | Nominal media share of the final volume.
Outer slope, C_target (1 - C_target/C_0) | 0.33333 | OD | Sensitivity of C_final to each step error: + for transfer, - for dilution (note 02, section 1, step 5).
u(eps_random) = CV | 0.035 | - | Type B, GUM 4.3.3: a CV is one standard deviation.
u(eps_systematic) = a / sqrt(3) | 0.046188 | - | Type B, rectangular distribution, GUM 4.3.7.
Covariance of the bias pair, r u(eps_sys)^2 | 0.0021333 | - | GUM 5.2.2 eq. 14 rearranged; shared bias via GUM F.1.2.3.
Input | Step | Dispense no. | Nominal volume (uL) | Weight w = v~/V~_step | u(x_i) | c_i (OD) | Signed contribution, c_i u(x_i) (OD) | Term in u_c^2 (OD^2) | Share of u_c^2
eps_random,t,1 | Transfer | 1 | 20 | 0.2 | 0.035 | 0.066667 | 0.0023333 | 5.4444e-06 | 7.143%
eps_random,t,2 | Transfer | 2 | 20 | 0.2 | 0.035 | 0.066667 | 0.0023333 | 5.4444e-06 | 7.143%
eps_random,t,3 | Transfer | 3 | 20 | 0.2 | 0.035 | 0.066667 | 0.0023333 | 5.4444e-06 | 7.143%
eps_random,t,4 | Transfer | 4 | 20 | 0.2 | 0.035 | 0.066667 | 0.0023333 | 5.4444e-06 | 7.143%
eps_random,t,5 | Transfer | 5 | 20 | 0.2 | 0.035 | 0.066667 | 0.0023333 | 5.4444e-06 | 7.143%
eps_random,t,6 | Transfer | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 0%
eps_random,t,7 | Transfer | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0%
eps_random,t,8 | Transfer | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 0%
eps_random,t,9 | Transfer | 9 | 0 | 0 | 0 | 0 | 0 | 0 | 0%
eps_random,t,10 | Transfer | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0%
eps_random,d,1 | Dilution | 1 | 20 | 0.4 | 0.035 | -0.13333 | -0.0046667 | 2.1778e-05 | 28.57%
eps_random,d,2 | Dilution | 2 | 20 | 0.4 | 0.035 | -0.13333 | -0.0046667 | 2.1778e-05 | 28.57%
eps_random,d,3 | Dilution | 3 | 10 | 0.2 | 0.035 | -0.066667 | -0.0023333 | 5.4444e-06 | 7.143%
eps_random,d,4 | Dilution | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0%
eps_random,d,5 | Dilution | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0%
eps_random,d,6 | Dilution | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 0%
eps_random,d,7 | Dilution | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0%
eps_random,d,8 | Dilution | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 0%
eps_random,d,9 | Dilution | 9 | 0 | 0 | 0 | 0 | 0 | 0 | 0%
eps_random,d,10 | Dilution | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0%
eps_systematic,t | Transfer bias |  |  |  | 0.046188 | 0.33333 | 0.015396 | 0.00023704 | 311%
eps_systematic,d | Dilution bias |  |  |  | 0.046188 | -0.33333 | -0.015396 | 0.00023704 | 311%
Covariance of the bias pair | 2 c_sys,t c_sys,d u(eps_sys,t, eps_sys,d) | [6 empty] | -0.00047407 | -622%
Total, u_c^2 | [7 empty] | 7.6222e-05
Combined standard uncertainty, u_c (OD) | [7 empty] | 0.0087305
Result (task 1.2: total and dominant terms)
Combined standard uncertainty, u_c (OD) | 0.0087305
Share of u_c^2: transfer random errors | 35.71%
Share of u_c^2: dilution random errors | 64.29%
Share of u_c^2: biases net (two variances + covariance) | 0%
Largest random-error term (share of u_c^2) | 28.57%
Input with the largest random-error term | eps_random,d,1